## Slide 1: Emission Measurements “a review”
Andy Griffin

## Slide 2
Objective
Make effective emission measurements [Radiated / Conducted, open site or chamber]
Repeatable measurements [No point in doing a measurement if you can not repeat it]
Use less chamber time [Chambers are very expensive, they need to be used effectively]
More relevant and accurate information from the time in the chamber ! [Attempt to eliminate possible sources of software, EUT and user errors]

## Slide 3
How do we do this ?
Quick Preview Method [why spend valuable time making measurements that have no relevance]
Interactive test, not reliant fully on software [software may not give you the correct answer, hence you can on rely on it to do certain functions]
Better understanding of the test procedure [Given processes needs to be well documented, followed plus training …etc ]
Better understanding of equipment operation [It helps to understand the operation of the equipment, then you can determine if it is doing the measurement correctly !]
Knowing the limitations of the test equipment [all equipment has limitations, knowledge of these will help you to perform better measurements, ie noise floor, overloading, frequency counters….. ]

## Slide 4
How do we do this ?
Quick Preview Method
  Use Spectrum Analyser Large Spans
  Max hold function
  Full rotation EUT
  Both polarities
  Antenna Heights ?
Warning !!!!
Loose both Frequency + Amplitude Accuracy

## Slide 5
Accuracy [of spectrum Analyzers]
Frequency
  Span frequency accuracy
  Span test time
Amplitude
  Peak Detector - worst case
  Pessimistic
Signal Resolution
  Span Signal Resolution
  Span Signal Resolution

## Slide 6
Specification Limit
If this was the entire profile why worry about signals greater than 10dB below the limit !

## Slide 7
Signals of interest
Only interested in the two signals Highlighted

## Slide 8
Resolution
Additional signal “within the peak” but why worry about it because it is well below the limit !

## Slide 9
Resolution
Example show data capture using 4 sub sweeps [trace 1] and no sub sweep [trace 2]

## Slide 10
Resolution
“Zoomed” in section, showing how data has not been lost but just not resolved.

## Slide 11
Details
Time [Don’t spend time investigating signals which are not relevant !]
Resolution [with a span of 400MHz and 400 points across the screen you can not resolve signal less the 1MHz ?]
Start / Stop / Split Frequency [May miss frequencies ?]
Need is a test receiver ? [They are generally easier to use with better preslection, noise floor, frequency accuracy but they are no good at spanning large frequency ranges, hence it is very difficult to make effective emission measurements with them !]

## Slide 12
Software Weakness
Time varying signals [If you have a time varying signal it is very difficult for software to maximise the signal]
Transient Signals [Transients from ambients, EUT, turntables can fool software into think there is a higher peak than there actually is !]
What is the correct frequency ? [Analyser frequency accuracy is generally 2% of Span, hence in a 100MHz that is 2MHz, how can it give you the correct frequency when you zoom in to take a closer look !]
Setting accuracy of the turntables/towers [Towers / Turntable have inertia and when the software states 45degrees, it may be at 41, plus they may move too quickly to resolve the signals]

## Slide 13
Software Weakness
Find peak routines [Find peak routines work quite well but there is a “trade off” between finding all the signals and finding enough, it is no good find 15 narrow band signals on top a broadband signal]
Narrow band signals [Narrow band signal may be “masked” by broadband signals and hence ignored, if this broadband signal is above the average limit then this may lead to serious errors]
Above 1GHz [above 1GHz very large frequency spans have to covered, hence the only effective way of performing the test is with user interaction and manual manipulation]

## Slide 14
Time Varying Signal ?
This is a time varying signal measured whilst rotation of the turntable, how is the software going to find the worst case signal level/azimuth.

## Slide 15
Time Varying Signal ?
This is a time varying signal measured whilst changing height of the antenna, how is the software going to find the worst case signal level/height.

## Slide 16
Frequency Accuracy
This highlights frequency accuracy problems of analysers in large spans. Trace 1, Trace 2 as previous defined and the “+” is the actual emission !

## Slide 17
Effective Measurements
Quick Preview Method
Interactive test, not reliant fully on software
Good test procedure
Better understanding of equipment operation
Knowing the limitations of the test equipment
  and……. with the use of VASONA, you too can make effective EMC measurements
  See the new presentation - Vasona -Helping You to Make Effective EMC Measurements